[image: image1.wmf] 

[image: image2.wmf] 

Version 3.0    written on 22/sep/2006


	Slot:vl04_a


List of Steps/Checks for powering up VELO-modules in the pit


Version 2.0   last edited on 20 May 2008

	Chief worker: Hans
	Date/time of start: 04/06/08 9:30

	Rest of team: Anne & experts
	Date/time of completion: 13:15

	Executive summary: The currents of both the R and Phi sensors are fluctuating, nevertheless we decided to take data with both HV on 100V.
The noise plots looked good. Bad channels as expected, but also a bad link on the R-sensor. Cable plots for the R sensor did not look so good and we did another delay scan. The offset found is the same as the offset found for VL02_A. (We’re going to preset this offset now for VL05_A). With the improved delay settings the cable plots for the R sensor were better. Although the settings were also changed for the Phi sensor (although the cabletest looked OK) the cabletest stayed quite the same.


!! Do not hesitate to add comments/notes in the margins !!

1. Slot label: vl04_av
2. Module number:  M63
3. Verify that warm cooling is already “ON” according to instructions on Twiki page.
4. Check that LV and HV of this module are “OFF” and tick the box indicating that LV and HV are indeed “OFF”:  X
5. Write down the numbers of the repeater boards

	
	
	Repeater board

	Top
	R
	21

	Bottom
	Phi
	84


6. Write down temperatures of the R-side
	
	cookie
	R-NTC0
	R-NTC1
	R-Volt.reg1
	R-Volt.reg2

	Temp.[oC]
	
	10.7
	11.1
	22.3
	22.2


             Write down temperatures of the Phi-side
	
	cookie
	Phi-NTC0
	Phi-NTC1
	Phi-Volt.reg1
	Phi-Volt.reg2

	Temp.[oC]
	
	10.7
	11.0
	23.0
	23.1


7. If cookie temperature gets above 14 OC, check cooling system according to instructions in Twiki. 
8. If NTC0-NTC1 >2 oC something is wrong.  Please contact a module/cooling expert.

9. Write down differential pressure DP411: -0.6 mbar

If  DP < -4.7 OR  DP> 1.7 mbar, you have to be careful because you are close to the limits when the Gas Injection System needs to be operational. Be prepared to switch “OFF” the LV.
10. Configure TELL1s from top panel and choose correct activity to see if TELL1 respond.
11. You are now going to power up the LV of the module, while keeping a close look at the temperatures and DP411. This is delicate job for at least two persons.
a. Switch on LV easy-board for both sides (6 channels)  
b. Switch on the LV at the control board for the R-side  of the module according to the instructions in Twiki.
c. Select all 16 Beetles of the R-side, and configure. Write down R-hybrid current: 3.59 (Expected values is  ~3.6 A).
d. Switch on the LV at the control board for the Phi-side of the module.

e. Select all 16 Beetles of the Phi-side, configure and wait 4 minutes. Write down R-NTC0: 21.6°C , R-NTC1: 27.4°C , Phi-NTC0: 21.5°C , Phi-NTC1: 26.1°C and phi-hybrid current: 3.58 A (Expected values are NTC0: 20-24 °C, NTC1: 25-29  0C,  and ~3.6 A)
12. IV-curve of the R-sensor:

a. Take an IV-curve according to the corresponding instructions.  Max voltage of 100 V.  Store the data in directory /group/velo/modulepowerup/IVscan.  
i. File name: 080604_1016_VL04_AT_Ch00_IVscan_10C.txt
b. Current at maxium voltage from Liverpool db is: 2.6 A.    

c. At maximum voltage, the leakage current at IP8 amounts to: 2.2-5.2 A.  (Note Liverpool data at 20 C, so expected ~2× more current at ~27 C)
13. IV-curve of the Phi-sensor:
a. Take an IV-curve according to the corresponding instructions.  Max voltage of 100 V.  Store the data in directory /group/velo/modulepowerup/IVscan.
i. File name 080604_1017_VL04_AB_Ch01_IVscan_10C.txt
b. Current at maximum voltage from Liverpool db is:  3.9 A
c. At maximum voltage, the leakage current at IP8 amounts to: 4.0-5.5 A.
(Note Liverpool data at 20 C, so expected ~2× more current at ~27 C)
d. Bias the R-sensor with 100 V and the Phi-sensor with 100 V. If you observe a deviation larger than 10% from the IV-data just taken, put HV to zero and consult a module expert.
14. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on. 
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)

i. File name: 080604_1107_VL04_A_modulepowerup_ntp_R100V_P100V.dat
b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File name: 080604_1112_VL04_A_modulepowerup_tp_R100V_P100V.dat
c. File name: Cable-test (Activity/Recipe: CABLETEST, 5000 events)

i. File name: 080604_1116_VL04_A_modulepowerup_cabletest_R100V_P100V.dat
ii. 080604_1228_VL04_A_modulepowerup_cabletest_R100V_P100V.dat
(file with new and improved delay settings)
15. Analyze the data according to the instructions. Print and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: x
b. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor : x
c. Any remarks about the data of the R-side ???

d. Tick the box if the data of the Phi-sensor look OK: x
e. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor : x
f. Any remarks about the data of the Phi-side ???
g. Tick the box if the cable-test is ok: first try not good, second try better
16. Dead/noisy channels of the R-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: 04-06-08
	

	
	285
	285
	
	

	
	560
	560
	560
	

	
	793
	793
	793
	

	
	795
	795
	795
	

	
	796
	796
	796
	

	
	845
	
	
	

	
	882
	882
	
	

	
	1121
	1121
	1121
	

	
	1123
	1123
	1123
	

	
	1124
	1124
	1124
	

	
	1297
	1297
	1297
	

	
	1299
	1299
	1299
	

	
	1301
	
	1301
	

	
	1348
	
	1348
	

	
	1353
	
	1353
	

	
	1679
	1679
	1679
	

	
	1680
	1680
	1680
	

	
	1681
	
	1681
	

	
	1682
	1682
	1682
	

	
	1683
	
	1683
	

	
	1684
	1684
	1684
	

	
	1685
	1685
	1685
	

	
	1686
	1686
	1686
	

	
	1726
	
	
	

	
	1807
	1807
	1807
	

	
	1808
	1808
	1808
	

	
	1810
	1810
	1810
	

	
	
	639
	639
	

	
	
	896
	896
	

	
	
	511
	511
	

	
	
	
	414
	1923/24/25

	
	
	
	Bad link 800-832
	1949


17. Dead/noisy channels of the Phi-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: dd-mm-yy
	

	
	623
	623
	623
	

	
	
	586
	
	

	
	
	1194
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	


18. Write down DP from DP411: -0.55 mbar .  Expected value between -5 mbar and +2 mbar.

19. Don’t forget to fill in the “Executive summary” and the “Date/time of completion”.
20. Copy your hand-written notes, numbers and ticked boxes to the digital version and store it at the following location: /group/velo/modulepowerup/checklists






PAGE  
page 1 of 4

_982663571.doc



