[image: image1.wmf] 

[image: image2.wmf] 

Version 3.0    written on 22/sep/2006


	Slot:VL09_C


List of Steps/Checks for powering up VELO-modules in the pit


Version 1.5   last edited on 8 May 2008

	Chief worker:MG
	Date/time of start: 13/5 14:30

	Rest of team: IM,BV, OB, MvB, ER
	Date/time of completion: 13/5 19:30

	Executive summary: Problems with configuring tellc24, solved 9/5/08

Module test all OK
Fluctuations in HV current understood as measurement precision of +/-0.8uA according to ISEG manual.


!! Do not hesitate to add comments/notes in the margins !!

1. Slot label: VL09_C
2. Module number:  44
3. Verify that warm cooling is already “ON” according to instructions on Twiki page.
4. Copy the dead/noisy channel lists of R and Phi-side as measured at Liverpool. from http://lbweb01.cern.ch/VELO/2 to the tables in point 21 and 22, respectively. 
5. Copy the dead/noisy channel list of R and Phi-side as measured at Assembly Lab to tables in point 21 and 22.  Located on pcveloassembly4.cern.ch in directory /scratch/lhcbvelo/assemblyData/Mxx.  File names modXX_checkout_150V_ntp_(r/p)_noisedata.txt
6. Find IV data from Liverpool db/Assembly lab elog.  Use Liverpool db measurement to set max voltage/expected bias current in points 16 and 17
7. Check that LV and HV of this module are “OFF” and tick the box indicating that LV and HV are indeed “OFF”:  X
8. Write down the numbers of the repeater boards

	
	
	Repeater board

	Top
	Phi
	32

	Bottom
	R
	2


9. Write down temperatures of the R-side
	
	cookie
	R-NTC0
	R-NTC1
	R-Volt.reg1
	R-Volt.reg2

	Temp.[oC]
	10.9
	10.5
	10.8
	20.8
	21.0


             Write down temperatures of the Phi-side
	
	cookie
	Phi-NTC0
	Phi-NTC1
	Phi-Volt.reg1
	Phi-Volt.reg2

	Temp.[oC]
	
	10.6
	11.2
	20.1
	20.2


10. If cookie temperature gets above 14 OC, check cooling system according to instructions in Twiki. 
11. If NTC0-NTC1 >2 oC something is wrong.  Please contact a module/cooling expert.

12. Write down differential pressure DP411:  0.08 mbar

If  DP < -4.7 OR  DP> 1.7 mbar, you have to be careful because you are close to the limits when the Gas Injection System needs to be operational. Be prepared to switch “OFF” the LV.
13. Configure TELL1s from top panel and choose correct activity to see if TELL1 respond.
14. You are now gradually going to power up the module, while keeping a close look at the temperatures and DP411. This is delicate job for at least two persons.
a. Switch on LV easy-board for both sides (6 channels)  
b. Switch on the LV at the control board for the R-side of the module according to the instructions in Twiki.
c. Select all even Beetle chips of R-side, configure and wait 4 minutes. Write down R-NTC0: 14.3°C, R-NTC1:16.8°C and hybrid current: 2.2 A (Expected values are R-NTC0:13-17 0C,  R-NTC1:16-19 0C, and ~2.2 A).
d. Select all 16 Beetles of the R-side, configure and wait 4 minutes. Write down R-NTC0: 17.4°C , R-NTC1: 20.0°C, hybrid current: 3.6 A (Expected values are R-NTC0:15-19 °C, R-NTC1:19-21°C and ~3.6 A).
e. Switch on the LV at the control board for the Phi-side of the module.

f. Select all even Beetle chips of Phi-side, configure and wait 4 minutes. Write down Phi-NTC0: 20.2°C , Phi-NTC1: 24.6°C, hybrid current: 2.2 A  (Expected values are Phi-NTC0:18-22°C, Phi-NTC1: 22-260C,   and ~2.2 A).
g. Select all 16 Beetles of the Phi-side, configure and wait 4 minutes. Write down Phi-NTC0: 21.8°C , Phi-NTC1: 26.9°C and hybrid current: 3.6A (Expected values are Phi-NTC0: 20-24 °C, Phi-NTC1: 25-29  0C,   and ~3.6 A)
15. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on.
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)
i. File Name:
080513_1650_VL09_C_modulepowerup_ntp_R0V_P0V.dat
b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File Name: 080513_1659_VL09_C_modulepowerup_tp_R0V_P0V.dat
16. Analyze the data according to the instructions in Twiki and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: x 

b. Tick the box if the data of the Phi-sensor look OK: x
c. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor :  x
d. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor :  x
17. IV-curve of the R-sensor:

a. Take an IV-curve according to the corresponding instructions.  Max voltage of  100 V which is the lower of: voltage at sensor draws 20 A or 100 V. Store the data in directory /group/velo/modulepowerup/IVscan.  
i. File name: 080513_1719_VL09_C_Ch03_Ivscan_10C.dat
b. Current at maxium voltage from Liverpool db is: 2.3 A.    

c. At maximum voltage, the leakage current at IP8 amounts to: 2A.
18. IV-curve of the Phi-sensor:
a. Take an IV-curve  according to the corresponding instructions.  Max voltage of 100 V which is the lower of: voltage at sensor draws 20 A or 100 V. Store the data in directory /group/velo/modulepowerup/IVscan.
i. File name:  080513_1721_VL09_C_Ch02_Ivscan_10C.dat
b. Current at maximum voltage from Liverpool db is:  1.7A
c. At maximum voltage, the leakage current at IP8 amounts to: 2 A.
d. Bias the R-sensor with 100 V and the Phi-sensor with 100 V. If you observe a deviation larger than 10% from the IV-data just taken, put HV to zero and consult a module expert.
19. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on.
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)

i. File name: 080513_1756_VL09_C_modulepowerup_ntp_R100V_P100V.dat
b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File name: 080513_1747_VL09_C_modulepowerup_tp_R100V_P100V.dat
c. File name: Cable-test (Activity/Recipe: CABLETEST, 5000 events)

i. File name: 080513_1801_VL09_C_modulepowerup_cabletest_R100V_P100V.dat
20. Analyze the data according to the instructions. Print and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: x
b. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor : x
c. Any remarks about the data of the R-side 
d. Tick the box if the data of the Phi-sensor look OK: x
e. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor : x
f. Any remarks about the data of the Phi-side 
g. Tick the box if the cable-test is ok: x 
21. Dead/noisy channels of the R-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: 06-05-08
	

	
	11
	1
	11
	

	
	12
	2
	12
	

	
	13
	3
	13
	

	
	14
	4
	14
	

	
	15
	5
	15
	

	
	18
	6
	19
	

	
	918
	7
	918
	

	
	919
	8
	919
	

	
	1241
	9
	1241
	

	
	1779
	10
	1779
	

	
	1780
	11
	1780
	

	
	1809
	12
	1809
	

	
	1810
	13
	1810
	

	
	1897
	14
	
	

	
	1922
	15
	1922
	

	
	2005
	18
	2005
	

	
	2040
	918
	2040
	

	
	2041
	919
	2041
	

	
	2042
	1241
	2042
	

	
	
	1631
	
	


	
	
	1779
	
	

	
	
	1780
	
	

	
	
	1809
	
	

	
	
	1810
	
	

	
	
	1922
	
	

	
	
	2005
	
	

	
	
	2040
	
	

	
	
	2041
	
	

	
	
	2042
	
	


22. Dead/noisy channels of the Phi-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: dd-mm-yy
	

	
	217
	217
	217
	

	
	308
	310
	308
	

	
	309
	311
	309
	

	
	310
	597
	310
	

	
	311
	1077
	311
	

	
	498
	1108
	
	

	
	597
	1187
	597
	

	
	1077
	1192
	1077
	

	
	1108
	1399
	1108
	

	
	1187
	1535
	1187
	

	
	1192
	1827
	1192
	

	
	1399
	1864
	1399
	

	
	1535
	1876
	1535
	

	
	1827
	
	1827
	

	
	1876
	
	1876
	


23. Write down DP from DP411: 0.07 mbar .  Expected value between -5 mbar-+2 mbar.

24. Don’t forget to fill in the “Executive summary” and the “Date/time of completion”.
25. Copy your hand-written notes, numbers and ticked boxes to the digital version and store it at the following location: /group/velo/modulepowerup/checklists






PAGE  
page 5 of 5

_982663571.doc



