[image: image11.wmf] 

[image: image12.wmf] 

Version 3.0    written on 22/sep/2006


	Slot:VL11_C


List of Steps/Checks for powering up VELO-modules in the pit


Version 1.5   last edited on 8 May 2008

	Chief worker:
	Date/time of start:080508-12:30

	Rest of team: TA, SB, HdV
	Date/time of completion:080508-12:50

	Executive summary: Current too phi hybrid too low (0.18 A per chip instead of nominal 0.21 A).   The Vcon is consistant with current being low with correct cable resistance.   Acts like incorrect initialization or less than 2.5 V out of repeater.   Data taken with 0V HV.   Noise on phi side 0.3 ADC, which is consistant with undervoltage.  HV scan is normal.



!! Do not hesitate to add comments/notes in the margins !!

1. Slot label: VL11_C
2. Module number:  55
3. Verify that warm cooling is already “ON” according to instructions on Twiki page.
4. Copy the dead/noisy channel lists of R and Phi-side as measured at Liverpool. from http://lbweb01.cern.ch/VELO/2 to the tables in point 20 (??) and 21 (??), respectively. 
5. Copy the dead/noisy channel list of R and Phi-side as measured at Assembly Lab to tables in point 20 (??) and 21 (??).  Located on pcveloassembly4.cern.ch in directory /scratch/lhcbvelo/assemblyData/Mxx.  File names modXX_checkout_150V_ntp_(r/p)_noisedata.txt
6. Find IV data from Liverpool db/Assembly lab elog.  Use Liverpool db measurement to set max voltage/expected bias current in points 16 and 17(??)
7. Check that LV and HV of this module are “OFF” and tick the box indicating that LV and HV are indeed “OFF”:  
[image: image1.wmf]W


8. Write down the numbers of the repeater boards

	
	
	Repeater board

	Top
	R/Phi
	24

	Bottom
	R/Phi
	36


9. Write down temperatures of the R-side
	
	cookie
	R-NTC0
	R-NTC1
	R-Volt.reg1
	R-Volt.reg2

	Temp.[oC]
	
	10.5
	10.8
	22.3
	22.3


             Write down temperatures of the Phi-side
	
	cookie
	Phi-NTC0
	Phi-NTC1
	Phi-Volt.reg1
	Phi-Volt.reg2

	Temp.[oC]
	
	10.4
	11.0
	22.1
	21.9


10. If cookie temperature gets above 14 OC, check cooling system according to instructions in Twiki. 
11. If NTC0-NTC1 >2 oC something is wrong.  Please contact a module/cooling expert.

12. Write down differential pressure DP411:  +0.9 mbar

If  DP < -4.7 OR  DP> 1.7 mbar, you have to be careful because you are close to the limits when the Gas Injection System needs to be operational. Be prepared to switch “OFF” the LV.
13. Configure TELL1s from top panel and choose correct activity to see if TELL1 respond.
14. You are now gradually going to power up the module, while keeping a close look at the temperatures and DP411. This is delicate job for at least two persons.
a. Switch on LV easy-board for both sides (6 channels)  
b. Switch on the LV at the control board for the R-side of the module according to the instructions in Twiki.
c. Select all even Beetle chips of R-side, configure and wait 4 minutes. Write down R-NTC0: 13.4°C, R-NTC1:15.4°C and hybrid current: 2.18 A (Expected values are R-NTC0:13-17 0C,  R-NTC1:16-19 0C, and ~2.2 A).
d. Select all 16 Beetles of the R-side, configure and wait 4 minutes. Write down R-NTC0: 16.6°C , R-NTC1: 18.7°C, hybrid current: 3.56 A (Expected values are R-NTC0:15-19 °C, R-NTC1:19-21°C and ~3.6 A).
e. Switch on the LV at the control board for the Phi-side of the module.

f. Select all even Beetle chips of Phi-side, configure and wait 4 minutes. Write down Phi-NTC0: 19°C , Phi-NTC1: 22.8°C, hybrid current:  1.35 A  (Expected values are Phi-NTC0:18-22°C, Phi-NTC1: 22-260C,   and ~2.2 A).
g. Select all 16 Beetles of the Phi-side, configure and wait 4 minutes. Write down Phi-NTC0: ____°C , Phi-NTC1: ____°C and hybrid current: 2.9A (Expected values are Phi-NTC0: 20-24 °C, Phi-NTC1: 25-29  0C,   and ~3.6 A)
15. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on.
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)
i. File Name: 080508_1241_VL11_C_modulepowerup_ntp_R0V_P0V
b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File Name:

080508_1246_VL11_C_modulepowerup_tp_R0V_P0V

16. Analyze the data according to the instructions in Twiki and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: 
[image: image2.wmf]W

 yes
b. Tick the box if the data of the Phi-sensor look OK: 
[image: image3.wmf]W

no, noise 0.3 ADC
c. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor :  
[image: image4.wmf]W

yes
d. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor :  
[image: image5.wmf]W

No power, no pulses
17. IV-curve of the R-sensor:

a. Take an IV-curve according to the corresponding instructions.  Max voltage of  100 V which is the lower of: voltage at sensor draws 20 A or 100 V. Store the data in directory /group/velo/modulepowerup/IVscan.  
i. File name:  080508_1136_VL11_C_CH06_IVScan_10C
b. Current at maxium voltage from Liverpool db is: 2.2 A.    

c. At maximum voltage, the leakage current at IP8 amounts to: 3.6A.
18. IV-curve of the Phi-sensor:
a. Take an IV-curve  according to the corresponding instructions.  Max voltage of 100 V which is the lower of: voltage at sensor draws 20 A or 100 V. Store the data in directory /group/velo/modulepowerup/IVscan.
i. File name:  080508_1136_VL11_C_CH07_IVScan_10C
b. Current at maximum voltage from Liverpool db is:  2.1 A
c. At maximum voltage, the leakage current at IP8 amounts to: 11.6 A.
d. Bias the R-sensor with 100 V and the Phi-sensor with 100 V. If you observe a deviation larger than 10% from the IV-data just taken, put HV to zero and consult a module expert.
19. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on.
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)

i. File name:

b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File name:

c. File name: Cable-test (Activity/Recipe: CABLETEST, 5000 events)

i. File name:
20. Analyze the data according to the instructions. Print and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: 
[image: image6.wmf]W


b. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor : 
[image: image7.wmf]W


c. Any remarks about the data of the R-side ???

d. Tick the box if the data of the Phi-sensor look OK: 
[image: image8.wmf]W


e. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor : 
[image: image9.wmf]W


f. Any remarks about the data of the Phi-side ???
g. Tick the box if the cable-test is ok: 
[image: image10.wmf]W

 
21. Dead/noisy channels of the R-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: 06-05-08
	

	
	734
	
	
	

	
	746
	
	
	

	
	770
	770
	
	

	
	1053
	1053
	
	

	
	1288
	
	
	

	
	1660
	1660
	
	

	
	1686
	1686
	
	

	
	1690
	1690
	
	

	
	
	896
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	


22. Dead/noisy channels of the Phi-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: dd-mm-yy
	

	
	573
	573
	
	

	
	799
	799
	
	

	
	875
	875
	
	

	
	1026
	1026
	
	

	
	1271
	1271
	
	

	
	1373
	1373
	
	

	
	1632
	1632
	
	

	
	1693
	1693
	
	

	
	1696
	1696
	
	

	
	1699
	1699
	
	

	
	1701
	1701
	
	

	
	1724
	1724
	
	

	
	1733
	1733
	
	

	
	1795
	1795
	
	

	
	1829
	1829
	
	

	
	
	42
	
	

	
	
	85
	
	

	
	
	117
	
	

	
	
	
	
	


23. Write down DP from DP411: _____ mbar .  Expected value between -5 mbar and +2 mbar.

24. Don’t forget to fill in the “Executive summary” and the “Date/time of completion”.
25. Copy your hand-written notes, numbers and ticked boxes to the digital version and store it at the following location: /group/velo/modulepowerup/checklists






PAGE  
page 3 of 5

_1269551616.unknown

_982663571.doc



