[image: image1.wmf] 

[image: image2.wmf] 

Version 3.0    written on 22/sep/2006


	Slot:vl16_a


List of Steps/Checks for powering up VELO-modules in the pit


Version 2.0   last edited on 20 May 2008

	Chief worker: HdV, MG, LD
	Date/time of start:080605 13.35

	Rest of team: 
	Date/time of completion:080605 14.56

	Executive summary: 

ITrip error on HV chan01 during ramp down after IV scan.  Acknowledging the error in the expert panel did not reset the error.
Ignored alarm and continued with normal operation.



!! Do not hesitate to add comments/notes in the margins !!

1. Slot label: vl16_a
2. Module number:  M21
3. Verify that warm cooling is already “ON” according to instructions on Twiki page.
4. Check that LV and HV of this module are “OFF” and tick the box indicating that LV and HV are indeed “OFF”:  X
5. Write down the numbers of the repeater boards

	
	
	Repeater board

	Top
	R
	51

	Bottom
	Phi
	76


6. Write down temperatures of the R-side
	
	cookie
	R-NTC0
	R-NTC1
	R-Volt.reg1
	R-Volt.reg2

	Temp.[oC]
	
	11.19
	11.73
	21.79
	21.05


             Write down temperatures of the Phi-side
	
	cookie
	Phi-NTC0
	Phi-NTC1
	Phi-Volt.reg1
	Phi-Volt.reg2

	Temp.[oC]
	
	11.13
	11.50
	23.07
	23.29


7. If cookie temperature gets above 14 OC, check cooling system according to instructions in Twiki. 
8. If NTC0-NTC1 >2 oC something is wrong.  Please contact a module/cooling expert.
9. Write down differential pressure DP411:  __-0.2__ mbar

If  DP < -4.7 OR  DP> 1.7 mbar, you have to be careful because you are close to the limits when the Gas Injection System needs to be operational. Be prepared to switch “OFF” the LV.
10. Configure TELL1s from top panel and choose correct activity to see if TELL1 respond.
11. You are now going to power up the LV of the module, while keeping a close look at the temperatures and DP411. This is delicate job for at least two persons.
a. Switch on LV easy-board for both sides (6 channels)  
b. Switch on the LV at the control board for the R-side  of the module according to the instructions in Twiki.
c. Select all 16 Beetles of the R-side, and configure. Write down R-hybrid current: _3.51__ A (Expected values is  ~3.6 A).
d. Switch on the LV at the control board for the Phi-side of the module.

e. Select all 16 Beetles of the Phi-side, configure and wait 4 minutes. Write down R-NTC0: __22.05__°C , R-NTC1: __27.80__°C , Phi-NTC0: _21.87___°C ,        Phi-NTC1: __26.56__°C and phi-hybrid current: __3.52__A (Expected values are NTC0: 20-24 °C, NTC1: 25-29  0C,  and ~3.6 A)
12. IV-curve of the R-sensor:

a. Take an IV-curve according to the corresponding instructions.  Max voltage of 100 V.  Store the data in directory /group/velo/modulepowerup/IVscan.  
i. File name: 080605_1353_VL16_AT_Ch01_IVscan_10C.txt
b. Current at maxium voltage from Liverpool db is: 2.1 A.    

c. At maximum voltage, the leakage current at IP8 amounts to: __3.8_A.  (Note Liverpool data at 20 C, so expected ~2× more current at ~27 C)
13. IV-curve of the Phi-sensor:
a. Take an IV-curve according to the corresponding instructions.  Max voltage of 100 V.  Store the data in directory /group/velo/modulepowerup/IVscan.
i. File name:  080605_1352_VL16_AB_Ch00_IVscan_10C.txt
b. Current at maximum voltage from Liverpool db is: 2.4 A
c. At maximum voltage, the leakage current at IP8 amounts to: _6__ A.
(Note Liverpool data at 20 C, so expected ~2× more current at ~27 C)
d. Bias the R-sensor with 100 V and the Phi-sensor with 100 V. If you observe a deviation larger than 10% from the IV-data just taken, put HV to zero and consult a module expert.
14. Collect events under each of the conditions below.
Remember to change the Beetle configuration each time, i.e. to change the recipe and switch them off/on.
a. Noise data, no TP (Activity/Recipe: PHYSICSNTP, 10000 events)

i. File name: 080605_1428_VL16_A_modulepowerup_ntp_R100_P100.dat
b. Noise data, with TP (Activity/Recipe: PHYSICSTP, 10000 events)
i. File name: 080605_1432_VL16_A_modulepowerup_tp_R100_P100.dat
c. File name: Cable-test (Activity/Recipe: CABLETEST, 5000 events)

i. File name: 080605_1436_VL16_A_modulepowerup_cabletest_R100_P100.dat
15. Analyze the data according to the instructions. Print and inspect the plots.

a. Tick the box if the data of the R-sensor look OK: 
b. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the R-sensor : X
c. Any remarks about the data of the R-side ??? Correlation plot broken
d. Tick the box if the data of the Phi-sensor look OK: X
e. Tick the box if the TP’s are where they are supposed to be, i.e. in channel 4 and 23 of the Phi-sensor : X
f. Any remarks about the data of the Phi-side ??? Correlation plot broken 
g. Tick the box if the cable-test is ok: X 
16. Dead/noisy channels of the R-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: 06-05-08
	

	
	74
	74
	74
	Dead

	
	376
	376
	376
	Dead

	
	400
	400
	400
	Dead

	
	420
	420
	420
	Dead

	
	424
	424
	424
	Dead

	
	444
	444
	444
	Dead

	
	468
	468
	468
	Dead

	
	484
	484
	484
	Dead

	
	502
	502
	502
	Dead

	
	537
	537
	537
	Dead

	
	795
	795
	795
	Dead

	
	832
	
	
	

	
	856
	856
	856
	Dead

	
	968
	968
	968
	Dead

	
	1023
	
	
	

	
	1062
	1062
	1062
	Dead

	
	1165
	1165
	1165
	Dead

	
	1380
	1380
	1380
	Dead

	
	1521
	
	1521
	Dead

	
	1531
	1531
	1531
	Dead

	
	1534
	1534
	1534
	Dead

	
	1631
	
	
	

	
	1633
	1633
	1633
	Noisy

	
	1634
	1634
	1634
	Noisy

	
	1641
	1641
	1641
	Dead

	
	1762
	1762
	1762
	Dead

	
	1950
	1950
	1950
	Dead

	
	2029
	2029
	2029
	Noisy

	
	2031
	2031
	2031
	Noisy

	
	2047
	
	
	

	
	
	896
	896
	Noisy

	
	
	1375
	1375
	Noisy

	
	
	1503
	1503
	

	
	
	1920
	1920
	Noisy

	
	
	
	
	

	
	
	
	
	


17. Dead/noisy channels of the Phi-sensor

a. Determine them according to the instructions

	
	Liverpool
	Assembly Lab.
	IP8 date: dd-mm-yy
	

	
	102
	
	102
	Dead

	
	103
	
	103
	Dead

	
	129
	
	129
	Dead

	
	377
	
	377
	Dead

	
	486
	
	486
	Dead

	
	578
	
	578
	Dead

	
	670
	
	670
	Dead

	
	870
	
	870
	Dead

	
	872
	
	872
	Dead

	
	1067
	
	1067
	Dead

	
	1082
	
	1082
	Dead

	
	1212
	
	1212
	Dead

	
	1218
	
	1218
	Dead

	
	1249
	
	1249
	Dead

	
	1353
	
	1353
	Dead

	
	1355
	
	1355
	Noisy

	
	1356
	
	1356
	Noisy

	
	1415
	
	1415
	Dead

	
	1704
	
	1704
	Dead

	
	1839
	
	1839
	Dead

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	


18. Write down DP from DP411: ___-0.3__ mbar .  Expected value between -5 mbar and +2 mbar.

19. Don’t forget to fill in the “Executive summary” and the “Date/time of completion”.
20. Copy your hand-written notes, numbers and ticked boxes to the digital version and store it at the following location: /group/velo/modulepowerup/checklists






PAGE  
page 1 of 6

_982663571.doc



